## Sheet1
Currently included:
index in calib file | Component | value (percent, k=1) | reference | L and E correlated | repeated meas. correlated | wavelengths correlated | comment
| random | stdev / sqrt(N) | empirical
4 | lamp | certificate / 2 | NMI | x | x | full | when using the same lamp
5 | lamp_age | 0.6*h/(50*sqrt(3)) | TR5 | x | x | filament | assuming systematic trend with age and both calibrations are done at the same time with the same lamp, current corr matrix assumed full
6 | lamp_power | u_current * 38 / wl | TR5 |  | x | filament | current corr matrix assumed full
7 | lamp_align | 0.1 | TR5 |  | x | full
| lamp_interp | 0 | contained in u_lamp |  | x
8 | panel | certificate / 2 | NMI |  | x | structured | current corr matrix assumed random
9 | panel_align | 0.1 | TR5 |  | x | structured | current corr matrix assumed random
10 | panel_interp | 0.1 | TR5 |  | x | full
11 | wl | 0.7 (3.3) nm | d_source/d_wl/sqrt(3) |  | x | structured | current corr matrix assumed random
12 | lab_stray | 0.1 | empirical |  | x | full
13 | panel_backrefl | 0.1 | empirical |  | x | full
14 | distance | 0.08(+0.4*0.8 for E) | TR5 |  | x | full
15 | alignment | 0.1 | TR5 |  | x | full
16 | temperature | 0.3 | guestimated |  | x | structured | currently a guess, will be updated. current corr matrix assumed random
17 | nonlinearity | wtd.p95 / 2 | empirical |  | x | structured | current corr matrix assumed random
18 | stray | 250-270 nm (0 for swir) | empirical | x | x | full | probably strongly correlated because L and E calibrations are done with spectrally rather similar sources, will be updated
19 | random in calibration | stdev / sqrt(N) | empirical |  | x |  | These are the random uncertainties in the gains, which will be in common to repeated measurements.
Currently missing:
| Component | value (percent, k=1) | reference | L and E correlated | repeated meas. correlated | wavelengths correlated | comment
| temporal interpolation uncertainty |  |  |  | structured | structured
| spectral interpolation uncertainty |  |  |  | structured | structured
| model error non-linearity | [5 empty] | included as uncertainty on gain?
model error combine VIS-SWIR
model error temp
model error stray
integration time uncertainty
windspeed unc
model error glint+foam
Not planning to include:
heterongeneity
clouds