## Sheet1
Currently included:
index in calib file | Component | value (percent, k=1) | reference | L and E correlated | repeated meas. correlated | VIS and SWIR correlated | wavelengths correlated | Correlation length | comment | actions
| random | stdev / sqrt(N) | empirical |  |  |  | random | 0
4 | lamp | certificate / 2 | NMI | x | x | x (complex) | full | inf | the same lap is used but its traceability is provided through different instruments, which would affect the error correlation structure. However the detailed correlation matrix is not provided with the correlation certificate.
5 | lamp_age | 0.6*h/(50*sqrt(3)) | TR5 | x | x | x (complex) | filament |  | assuming systematic trend with age and both calibrations are done at the same time with the same lamp, current corr matrix assumed full
6 | lamp_power | u_current * 38 / wl | TR5 |  | x | x (complex) | filament |  | current corr matrix assumed full
7 | lamp_align | 0.1 | TR5 |  | x | x | full | inf
| lamp_interp | 0 | contained in u_lamp |  | x
8 | panel | certificate / 2 | NMI |  | x | x (complex) | structured |  | current corr matrix assumed random
9 | panel_align | 0.1 | TR5 |  | x | x | structured |  |  | CHANGE to FULL
10 | panel_interp | 0.1 | TR5 |  | x |  | full | inf |  | not correlated
11 | wl | 0.7 (3.3) nm | d_source/d_wl/sqrt(3) |  | x |  | structured |  | current corr matrix assumed random
12 | lab_stray | 0.1 | empirical |  | x | x (complex) | full | inf
13 | panel_backrefl | 0.1 | empirical |  | x | x | full | inf
14 | distance | 0.08(+0.4*0.8 for E) | TR5 |  | x | x | full | inf
15 | alignment | 0.1 | TR5 |  | x | x | full | inf
16 | temperature | 0.3 | guestimated |  | x |  | structured |  | currently a guess, will be updated. current corr matrix assumed random
17 | nonlinearity | wtd.p95 / 2 | empirical |  | x |  | structured |  | current corr matrix assumed random
18 | stray | 250-270 nm (0 for swir) | empirical | x | x |  | full | inf | probably strongly correlated because L and E calibrations are done with spectrally rather similar sources, will be updated
19 | random in calibration | stdev / sqrt(N) | empirical |  | x |  | random | 0 | These are the random uncertainties in the gains, which will be in common to repeated measurements.
Currently missing:
| Component | value (percent, k=1) | reference | L and E correlated | repeated meas. correlated | VIS and SWIR correlated | wavelengths correlated |  | comment
| temporal interpolation uncertainty |  |  |  | structured |  | structured
| spectral interpolation uncertainty |  |  |  | structured |  | structured
| model error non-linearity | [7 empty] | included as uncertainty on gain?
model error combine VIS-SWIR
model error temp
model error stray
| integration time uncertainty | negligible
POLARISATION
calibration drift
windspeed unc
model error glint+foam
Not planning to include:
heterongeneity
clouds